On-Device Spike Classification for Wireless Neural Recording on a Commercial Bluetooth Low Energy SoC

Fredrik Willumsen Haug

Bioengineering

Dr. Saehyuck Oh and Professor Gage Hills

Table of Contents

Executive Summary

DEFINE

- Background and motivation

  - Background and Motivation

  - Stakeholders, Client, and End User of the Solution

  - Science, Technology, and Society considerations (STS)

- Existing Solutions and Previous Work

- Problem Statement

- Design Requirements / Technical Specifications

  - Description

    - Values

    - Justification

    - Measurement Method

  - Summary Table

DESIGN

- Consideration of Multiple Design Approaches

REFERENCES AND APPENDICES

- References

- Appendices

  - Preliminary Project Milestones and Budget

Project Executive Summary (150 words or less)

This thesis addresses a bandwidth bottleneck in wireless neural recording. Bluetooth Low Energy (BLE) 6.0, the leading low-power wireless protocol, has a 4 Mbps theoretical ceiling but reaches only ~1 Mbps in practice. Extracellular action potentials contain frequencies from around 300 Hz to 6 kHz and are sampled at 30 kHz as the standard in the field. at 16-bit resolution, each channel produces 480 kbps. BLE can therefore stream raw data from only two channels which is far below modern high-density probes with hundreds to thousands of channels. Instead of transmitting raw data for offline spike sorting, this project detects and classifies spikes on-device, transmitting only a unit ID and timestamp per spike. This cuts the per-spike data volume by orders of magnitude, enabling far more channels to be streamed wirelessly on a commercial BLE system-on-chip which is a capability not previously demonstrated on commercial hardware.

- Background and Motivation

  - Background and Motivation

Whether the substance of thought is reducible to physical processes is an unanswered question. Nevertheless, evidence clearly suggests that the organized interplay between neurons in the central nervous system gives rise to human sensory and cognitive capacity.1 The first plot of what is now known as the action potential – the voltage across an electrically excitable cell’s membrane as a function of time – was presented in 1868 by Professor Julius Brenstein.2 Brenstein’s result, derived from the tissue of a frog leg using his own invented differential rheotome, served as pivotal evidence for theories propounded by physicists and physiologists preceding Brenstein.2 Understanding the organized interplay between neurons in the central nervous system can be attempted through measuring the action potentials elicited by nerve cells constituting populations of neurons within, and, or across anatomically defined regions of the brain.3

In 2017, the Neuropixels silicon recording probe was presented. The Neuropixels 1.0 silicon probe is an intracranial, extracellular electrophysiology recording probe designed for rodent neural recordings, allowing for simultaneous recording of neurons from 384 recording channels.3 Since then, Neuropixels have released the second iteration of the probe, the Neuropixels 2.0, with a channel count up to 5120 recording sites.4 These recording systems have enabled groundbreaking discoveries in systems neuroscience.5 However, a higher channel count sets a higher bandwidth requirement, presenting a challenge.

Continuous sampling of neurons using high channel count silicon probes generates large amounts of data. The industry standard for sampling raw extracellular voltage is the Intan technologies analog frontend which uses a 16-bit ADC typically configured with a 30KHz sampling rate.6-7 Thus, the data rate for continuous recordings with the Neuropixels 2.0 quad base becomes 2.56 Gbps (see appendix for calculation). The current state of the art Bluetooth Low Energy (BLE) protocol has a theoretical maximum transmission capacity of 4 Mbps, which makes transmission of uncompressed raw data impossible with the power constraints needed for biomedical devices.8-9 Fortunately, for many questions of interest to neuroscientific inquiry, only action potential signatures are required, which constitutes only a small fraction of the continuous raw signal.10-11

  - Stakeholders, Client, and End User of the Solution

The end user of the solution are system neuroscientists performing in vivo behavioral electrophysiology recordings in rodents.

  - Science, Technology, and Society Considerations (optional for pre-proposal)

pass

- Existing Solutions and Previous Work

Among the most relevant prior art is the Lukito paper and its references.

- Problem Statement

The Bluetooth Low Energy 6.0 protocol’s bandwidth is too low to transmit raw extracellular electrophysiology data from more than two recording channels12 for offline spike sorting. [from rubric: … and lists the requirements of a model solution that addresses the shortfalls of existing/previous solutions.]

- Design Requirements / Technical Specifications

Need Resulting from Problem Statement

Transmit only neuron unit IDs (after classification) with corresponding timestamp values.

Functional Requirements

- Acquire simulated extracellular electrophysiology data from at least three recording channels

- Detect neural spike events from the acquired data

- Save the time at which each detected spike occurs

- Classify each detected spike and assign it a unit ID

- Transmit the unit IDs, each with a corresponding time stamp

System-Level Requirements

S1. No loss / keep up

No detected spike or the raw sample that produced it is dropped due to processing or transmission backlog up to a sustained throughput.

Technical Requirements

- Acquire simulated extracellular electrophysiology data from at least three recording channels

  - There must be at least three analog channels amplified and digitized through an analog-to-digital converter (ADC)

  - The ADC must sample fast enough to capture the frequency spectrum contained in a neural spike

  - The total digitization throughput must be high enough to sample every channel at its required per-channel rate

  - The digitized ADC code must preserve spike amplitude above the noise floor with enough bits

  - There must be a communication protocol in place with high enough throughput to get the ADC codes transmitted to a microprocessor without data loss

- Detect neural spike events

  - Neural spikes must be distinguished from the local field potential

- Save the time at which each detected spike occurs

  - The detection timestamp of each spike must be recorded

- Classify each detected spike and assign it a unit ID

  - Each spike must be classified according to some classification inference method

  - (I just realized that the offline machine/deep learning model is also a part of the system in a sense – not sure how this fits into the functional and technical requirements and technical specifications)

  - A unit ID must be assigned to the classified spike according to the pre-sorted library

- Transmit the unit IDs, each with a corresponding timestamp

  - Combine assigned unit ID and corresponding timestamp and transmit it over BLE with high enough total throughput to transmit all data from at least three channels without data loss

Technical Specifications

Justification

Measurement Method

  - Summary Table

[Insert a table that lists things in a structured and easily interpretable way here]

DESIGN

- Consideration of Multiple Design Approaches

  - Calibration Recording + Online Classification

One way to achieve this is to perform calibration recording first to have an assumed ground truth library for units for the subject.

  - Online Sorting

There is also the Osort algorithm which can be implemented. The problem with this is that we do not have all the data up front to perform clustering on which to my understanding would produce less reliable, sorted units compared to the offline sorting with supervised learning inference.

REFERENCES AND APPENDICES

- References

  - First reference

  - Second reference

  - …

  - Nth reference

- Appendices

  - Preliminary Project Milestones and Budget

My goal is to have a minimally viable product implementing the simplest form of classification algorithm by October 1st.